## Slide 1
J. Morel, N. Castagna
14IND13 Photind WP3: Development of fibre optic
measuring instruments and artefacts

## Slide 2: Main goals of WP3
Activities are divided into three tasks, with the main goals to develop:
Task 3.1 A fully traceable instrument for the measurement of the angular resolved modal distribution in multimode fibres, i.e. encircled angular flux (EAF).
Task 3.2 New artefacts and calibration techniques for the calibration of high resolution optical reflectometers (OTDR, OLCR) and of optical fibre sensors for distributed temperature measurements
Task 3.3 A novel portable absolute standard fibre detector based on nanotubes at cryogenic temperature.

## Slide 3: Task 3.1 Encircled Angular Flux
Guided modes in a multimode fibre experience different path lengths and losses. Correct measurements of quantities like insertion losses or bandwidth strongly depend on how different modes are populated
Modal distribution metrics like the Encircled Flux (EF) already exists and is already fully traceable but it is not applicable to step index large core multimode fibres
The Encircled Angular Flux (EAF) metrics overcomes this issue
Need for a fully traceable EAF measuring instrument

## Slide 4: Task 3.1 Encircled Angular Flux measurements
This metrics allows quantifying the modal distribution of asymmetrical intensity profiles, as those potentially observable in large core, step index multimode fibres
'

EAF is derived from the analysis of the far field pattern at the output of the MM fibre.
Corresponds to the relative power observed in a circle defined by each angle q’

## Slide 5: A3.1.1 Traceable EAF measuring instruments
Two fully independent EAF measuring instruments were developed by Arden and by METAS
  Different camera types (CCD, CMOS, different chip and pixel sizes, micro lens arrays)
  Image processing and data analysis programmed in Labview, using two completely different software

## Slide 6: A3.1 Arden-METAS first intercomparison
By using a common source, a first intercomparison has been realised between the two setup (november 2016)
Results presented in WP4

## Slide 7: A3.1.5 Uncertainty budget
Systematic or Statistic | Max Uncertainty (k=2) on EAF
Camera uniformity | 0.001 (0.3% @ q = 2°)
Camera non-linearity | 0.0006
Pixel resolution uncertainty | < 0.0001
Distance fibre-camera chip | 0.001
Correction of the refraction on the camera front window | 0.001
Repeatability (n=16) | 0.0001
Utot (k=2) | 0.0018
Uncertainty budget for EAF measurements fully established
Uncertainty calculation implemented in the METAS EAF measurement software

## Slide 8: A3.1.5 Uncertainty budget
Example 1 : uncertainty due to the measurement of DF
Camera
Micrometer Stage
DF

r
DF
EAF
EAF

## Slide 9: A3.1.5 Uncertainty budget
Example 2: Uncertainty due to the correction of the refraction in the camera front window

r
DF
DF
chip
source
EAF
theta (degree)
window
  Comparison between EAF measurements performed without and with window but corrected for the beam refraction shows comparable results

## Slide 10
PhotInd
Initial work has been with 50/125 and 62.5/125µm graded index fiber
Desire to expand measurement capability to as many different fiber types as possible
Created a pair of reference light sources with 50/125µm Step Index fiber
Illumination from 850nm LED
EAF in 50 / 125­µm Step Index Fiber
10
EMPIR JRP 14IND13 PhotInd 27th month meeting / 4-5.10.2017 / Bern

## Slide 11
PhotInd
Comparison with Graded Index
Chart: EAF in Step Index and Graded Index Fiber
11
EMPIR JRP 14IND13 PhotInd 27th month meeting / 4-5.10.2017 / Bern

## Slide 12: Status of activities in Task 3.1
A3.1.1 | Arden, supported by METAS, will design and build a system for traceable EAF measurements. The target accuracy will be specified considering the requirements fixed by the EAF limit template values provided by the industry or by normalisation bodies. | Arden, METAS On time | M24
A3.1.2 | METAS, supported by Arden, will perform in depth analysis of the performance of the system developed in A3.1.1, including optics, 2D optical detection, image and data processing. | METAS, Arden On time | M32
A3.1.3 | METAS and Arden will evaluate the impact of the proposed solution developed in A3.1.1 and A3.1.2 on relevant normalisation documents by discussing the achieved results with the corresponding Working Groups within the IEC (e.g. IEC TC86 WG4). | METAS, Arden Ongoing. First contacts already taken | M34
A3.1.4 | Modelling of the modal distribution in multimode fibres as the function of the light source parameters and coupling geometry, and analysis of mode dependent coupling losses will be carried out. JCM will develop interfaces to efficiently and accurately simulate the modal distribution and mode profiles in multimode fibres. These data will be used to analyse the influence of modal properties on coupling efficiency. METAS and Arden will provide specifications of multimode fibre setups and will test the developed interfaces. | JCM, METAS, Arden Finished | M30
A3.1.5 | Arden and METAS will write a report and for the system developed in A3.1.1 and A3.1.2. An uncertainty estimation will be included. Arden will send the report and uncertainty budget to the coordinator for submission to EURAMET as D6. | Arden, METAS Ongoing | M34

## Slide 13: Task 3.1 Next steps
Realisation of larger scale inter-comparisons (possibly within IEC)
Measurement of the EAF on Step Index fibres (larger core)
Measurements extended to other kind of fibres (POF in the visible)

## Slide 14: 3.2 Development of calibration artefacts
The aim of this task is:
To develop artefacts for distance scale calibration of optical reflectometers at the cm level and below and based on miniaturised optical recirculating delay lines allowing to produce evenly distributed and traceable reflection events.
To develop an artefact for the calibration of the attenuation scale of multimode reflectometers, based on modal distribution insensitive spectral attenuation reference fibres
To develop calibration techniques for optical fibre-based distributed temperature and vibration sensors
MC
a(l)
Modal control
Att. standard
Refl.

## Slide 15: A3.2.1 and A3.2.2 Fibre-coupled resonators for the calibration of the distance scale of optical reflectometers
Two different kinds of standards are currently evaluated
Integrated optics micro ring resonators of a length ranging from 60 mm to 50 mm have already been fabricated by WWU. These artefacts should allow achieving a distance resolution in 30 um up to 25 mm range and are intended for the calibration of the distance scale of Optical Low Coherence Reflectometers (OLCR)
Optical delay line made of standard optics components (made by METAS) of a length between 20 and 60 cm. The artefact will be used for the calibration of high resolution singlemode and multimode OTDR

## Slide 16: A3.2.1 and A3.2.2 Fibre-coupled micro ring resonators
Grating couplers
for vertical light coupling using SMF fibre arrays
A chip containing a set of 210 different ring resonators was fabricated
7 different fibre loop lengths
5 different gap distances (coupling ratio)
6 different coupler grating periods
SEM image of fabricated device

## Slide 17: A3.2.1 and A3.2.2 Fibre-coupled micro ring resonators
Fibre array
Chip
Setup for positioning the fibre array on the selected resonator ready:
The chip lays on a XY stage
The fiber array is fixed on a Z+f qy stage
Prototype calibration artefact ready for first measurements

## Slide 18: A3.2.1 and A3.2.2 First measurements with the micro ring resonators
L=9.72 mm
Gap=2.2 um
Grating period 1.11 um (l=1550 nm)
ECDL
PD
L
Spectral transmission
High sensitivity to the polarization
High coupling losses

## Slide 19: A3.2.3 Calibration artefact for the calibration of the attenuation scale of multimode OTDR
Goal: to develop a spectral attenuation reference fibre with controlled modal distribution
Evaluation of the repeatability of spectral attenuation measurements on a 62.5 mm fibre spool with controlled modal distribution successfully performed
Prototype ready for further in-depth evaluation with different OTDR
MC
a(l)
Modal control
Att. standard
OTDR
UA < 0.009 dB/km (k=2)

## Slide 20: A3.2.5 OTDR temperature measurements
Patent Application ES P201530793
  Photonic differentiation detector to determine both phase and amplitude of the backscattered pulse. High resolution and sensitivity measurements without reference state.
New results: OTDR calibration setup
  Simultaneous calibration of measurement error and resolution.
  Experimentally demonstrated on a commercial DTS system.
Publications
  Presented in SINFOTON’16 and 6th Spanish Conference on Metrology.
  Submitted to Óptica pura y aplicada (OPA).

## Slide 21: A3.2.6 Bragg grating sensors
Development of Bragg grating interrogator calibration techniques
Publications
  Presented in SINFOTON’16 and 6th Spanish Conference on Metrology.
  Submitted to Review of Scientific Instruments.

## Slide 22: A3.2.7 OTDR vibration measurements
Patent Application: ES P201531736
  Chirped pulses in coherent OTDR for vibration measurements.
  Single pulse characterization with high resolution and sensibility.
Next step: Reference systems and calibration techniques for cOTDR

## Slide 23: Status of activities in Task 3.2
A3.2.1 | WWU, with optical metrology support from METAS, will design a fibre‑coupled waveguide micro-resonator for the calibration of the distance scale of high resolution OTDR with cm and/or mm resolution. | WWU, METAS On time | M12
A3.2.2 | METAS will fabricate a distance scale calibration artefact, with support from WWU on the waveguide interconnection aspects, and will evaluate its performance through a series of validations by comparison to different reference instruments at METAS. An evaluation of its applicability will be carried out by various companies fabricating OTDRs (Luciol and JDSU have shown a strong interest in participating in this comparison. EXFO may also be interested). | METAS, WWU Ongoing | M30
A3.2.3 | METAS will design and fabricate a traceable calibration artefact for the calibration of the attenuation scale of multimode OTDR and will evaluate its performance through a series of validations by comparison to different reference instruments performed jointly with NPL, CSIC and various interested companies fabricating OTDRs. | METAS, NPL, CSIC Ongoing, according to schedule | M32
A3.2.4 | METAS and NPL will evaluate the impact of the solutions developed in A3.2.2 and A3.2.3 on relevant normalisation documents (IEC calibration standards). | METAS, NPL | M34
A3.2.5 | CSIC will develop reference systems and calibration techniques for temperature measurements in OTDR distributed optical-fibre sensors. NPL will contribute with the characterisation of OTDR. | CSIC, NPL Finished | M12
A3.2.6 | CSIC will develop reference systems and calibration techniques for temperature measurements in Bragg-grating-based quasi-distributed optical-fibre sensors. | CSIC Ongoing | M24
A3.2.7 | CSIC will develop reference systems and calibration techniques for vibration measurements in distributed OTDR optical-fibre sensors. | CSIC Ongoing | M34
A3.2.8 | METAS, WWU, NPL and CSIC will write a report on the fully traceable artefacts for the calibration of the distance scale of high resolution OTDR (A3.2.2) and for the attenuation scale of multimode OTDR (A3.2.3). An uncertainty estimation will be included. METAS will send the report to the coordinator for submission to EURAMET as D7. | METAS, WWU, NPL, CSIC | M34

## Slide 24: Task 3.2 Next steps
Calibration of the optical length of the micro ring resonators Optimization of the proposec concept
Evaluation of macro ring resonators based on classical optical components for the calibration of high resolution OTDR
Continuation of the activities on distributed sensors according to schedule

## Slide 26: A.3.3.2 Fibre coupling unit for primary standard detector
Quantity | Value
PDL Out 1 | < 0.04 dB
PDL Out 2 | < 0.06 dB
Switching repeatability Out 1 | < 0.0004 dB
Switching repeatability Out 2 | < 0.0009 dB
Splitting ratio Out 2 / Out 1 | 1.0116 ± 0.0006
Switch specs @ 1550 nm
APC
PC
PC
In
Out 1
Out 2
MEMS-based switch system built and fully validated
Switching capability: manual, TTL-trigger or USB

## Slide 27
Shortenning of the traceability chain
PhotInd CSIC January 2017
200 nm
900 nm
1600 nm
Trap detector
Si detector
NIR standard radiometer (InGaAs/Ge)
0.01%
0.20 %
n ..laser lines
spectrally independent (pyro)
0.50 %
CR
0.90 %
Direct comparison with thi fibre coubled absolute standard
0.30 to 0,5 %

## Slide 28: T3.3 Management
A3.3.1 delayed due to 9 months delayed procurement of a dedicated desktop cryo cooler in CMI partially caused by the fact that CMI requirements on temperature stability required more iterations with the producer.
The current month delay of A3.3.1 caused by it was partially compensated by the in-advance works on A3.3.2 completed thanks to the experimental works done in NIST in collaboration (below). We are then confident that the functioning system will be delivered by the deadline for A3.3.3 as expected.

## Slide 29: M9-M18 Activities
Portable mechanically cooling cryostat procured
  Critical parameter is very low vibration amplitude
  Cooling down cycle less then 4 h RT to 4K
  2-pole Rc thermodynamic filter implemented additionally to dump down fundamental cooling frequence
Custom designed and cooled chamber built.
Chamber is designed for 2 independent chips and independent fibers for each chip
All acivities are going on in close collaboration with Sources and Detectors Group , NIST Boulder
A3.3.1 delayed due to 9 months delayed procurement of a dedicated desktop cryo cooler in CMI partially caused by the fact that CMI requirements on temperature stability required more iterations with the producer.
The current month delay of A3.3.1 caused by it was partially compensated by the in-advance works on A3.3.2 completed thanks to the experimental works done in NIST in collaboration (below). We are then confident that the functioning system will be delivered by the deadline for A3.3.3 as expected.

## Slide 30: M9-M18 Activities
1. CMI facility for operating planar bolometric radiometer is ready:
Current sources for reference block and cavity
High metrological grade DC bridge for temperature sensors readings ready and characterised
2. CMI designed and developed the software to:
Control and stabilise the temperature of reference block and optical sensor
Automated procedures of static and dynamic substitution of ESR
Provides the traceable measurement and evaluation of ESR-based optical power
Implemented thermal link calibration of system
A3.3.2 Activities completed
All acivities are going on in close collaboration with Sources and Detectors Group , NIST Boulder

## Slide 31: M9-M18 Activities
Measurements ran on free space beam
CMI software positively tested in terms of integrity and functionality
PID optimased in term of stability of the reference block
Optical power measurement procedure and electrical power measurement procedure fully validated.
First real free space beam optical power measurement at 633 nm where performed
First testing measurements ran through in NIST Boulder in October 2016:
All acivities are going on in close collaboration with Sources and Detectors Group , NIST Boulder
VACNT Absorber
Heat-link

## Slide 32: M18-M27 Activities
Wire bonded Bolometer chip inserted in the housing
All thermal filter components had to be cleaned to remove vacuum oil residuals
Optical fiber coupled with the bolometer
Series of temperature and vacuum cycles performed to evaluate robustness of the system
Electrical wiring completed and tested before and after the cycles:
  Sensor temperature probe
  Sensor heating
  Reference block temperature sensor
  Reference block heating
Implemented Active cavity routine to implement dynamic substitution :
  A parallel PID process keeps the temperature of the bolometer constant. The difference between the power that the PID uses with and without illuminating the sensor gives the applied optical power
Assembling the system in CMI
All acivities are going on in close collaboration with Sources and Detectors Group , NIST Boulder

## Slide 33: Activities in WP3
Task 3.3 Novel fibre coupled power detector
A3.3.1 | CMI will integrate a carbon nanotube based chip into a suitable cryostat convenient for portable operation and will provide the electronic readout instrumentation. | CMI Ongoing | M9 Delayed
A3.3.2 | METAS, with collaboration from CMI, will design, manufacture and characterise the fibre coupling for the radiometer developed in A3.3.1. METAS will concentrate on the structure and on the optical properties of the fibre coupling system and CMI will work on the related thermal conductivity issues. | METAS, CMI Ongoing | M18 Development of switch: finished
 |  |  | Integration in radiometer: Delayed
A3.3.3 | CMI will assemble the prototype developed in A3.3.1 and A3.3.2 and check the functionality in terms of its sensitivity. | CMI | M30
A3.3.4 | CMI and METAS will perform the comparison of the prototype assembled in A3.3.3 with the currently established highest standard radiometer. | CMI, METAS | M32
A3.3.5 | CMI and METAS will write a report on the development and performance evaluation of the prototype fibre coupled carbon nanotube based cryogenic radiometer developed in A3.3.1 to A3.3.3 and evaluated in A3.3.4, and will send it to the coordinator for submission to EURAMET as D8. | CMI, METAS | M34

## Slide 34: Deliverables
The activities carried out in WP3 lead to following deliverables:
D6: Report on the demonstration of a fully traceable system for EAF measurements. Delivery date : M34.
D7: Report on the demonstration and validation, by inter-comparison, of fully traceable artefacts for the calibration of the distance scale of high resolution OTDR and for the attenuation scale of multimode OTDR. Delivery date : M34.
D8: Report on the development and performance evaluation of a prototype fibre coupled carbon nanotube based on cryogenic radiometer, with a target accuracy better than 0.5 % and the capability to measure optical power at 1550 nm with a relative uncertainty of less than 0.25 %. Delivery date : M34.

## Slide 35: Thank you very much for your attention

## Slide 36
Camera
Hamamatsu ORCA Flash 4.0LT
CMOS sensor
2048 x 2048 pixels
6.5 um pixel size
(13.3 x 13.3) mm effective area
AD conversion: 16 bit
Optomechanics
XYZ- alignment of the fibre
XYZ alignment of the camera
Implementable intensity filters
Light shield over the whole system
Camera
Fibre positioning unit (XYZ-)
Camera XYZ translation stage
Fibreand Filter holder
A3.1.1 METAS EAF Measuring instrument

## Slide 37
Micrometer
Translation Stage
Camera
Reference EF-compliant light source
Shielding from
External light
50µm or 62.5µm
MM fibre
Insert sub-header here
Camera
Micrometer Stage
Reference Light Source
Camera
Lumenera Lm11059
CCD sensor
(35x25) mm sensor, 9.0µm pixel size
A/D conversion: 14 bits
Canon EF Mount (44.0mm sensor to front of camera)
A3.1.1 Arden EAF Measuring instrument
37
EMPIR JRP 14IND13 PhotInd 27th month meeting / 4-5.10.2017 / Bern

## Slide 38: 3.1 Modelling of the modal distribution in MM fibres
These tools are now available for future EAF simulations in multimode fibre systems
LP11
Calculate the eigenmodes of a fibre
Modelling of a source
Modal distribution, depending on the coupling geometry
Simulate the modal propagation in a uniform fibre
Calculate overlap integrals
Finite elements method using JCMSolve interfaced with MatLab (Octave)
EMPIR JRP 14IND13 PhotInd 27th month meeting / 4-5.10.2017 / Bern